| Reported by (Name): | Geoffrey S. Ibbott, Ph.D |
| --- | --- |
| Organization: | International Electrotechnical Commission |
| Position Title: | Chair, Subcommittee 62C. Member, Working Group 1 |
| Activity: | Working Group meeting |
| Meeting Dates: | July 10 – 12, 2006 |
| Meeting Location: | Ottawa, Canada |
| Payment $: | Travel reimbursement |
| Reasons for Attending or not Attending | Attended to contribute to the writing of international standards for radiotherapy equipment. |
| Issues from Previous Meetings or Year: | See below. |
| General Description of Activities of the Organization and/or Meeting: | Working Group 1 IEC Sub-Committee 62C July 10-12, 2006, Ottawa Participants The meeting was attended by Lillicrap (convenor), Sethi (UK); Seidel (secretary), Höppner, Lehmann and Goldstein (Germany); Dvorak, Gerig and Simons (Canada); Ligthart (Netherlands); Johnsen, Marles, Lukey, Coon and Ibbott (US). Agenda 60601-2-29 Simulator safety standard Ligthart had distributed the document as a CD and national committee comments were discussed. Almost all were editorial, except one regarding leakage current. As the draft is in agreement with the general standard, the accelerator standard, and the previous edition of this standard, it was agreed to reject the comment. Another comment on gantry rotation speed was accepted. The remaining comments were editorial and will be handled by Guido. Guido will prepare another version for a CDV. Lillicrap does not want Ligthart to spend the time converting this draft to the new template for particular standards. 60601-2-17 Afterloader safety standard Ibbott agreed to finalize a CD of 60601-2-17 for the next meeting, using the new template. 60976 Accelerator performance standard Seidel had distributed the compilation of comments shortly after the deadline of July 7. The result of the vote was assumed to be positive, based on reports from attending members. Extensive discussion was held on the comments. The recommendations for changes to clause 8.1.6 were discussed extensively, and it was agreed that the principal drafters would consider the AAPM and European publications and prepare another alternate proposal for discussion prior to the Delft meeting. It is hoped that following the discussion in Delft, an FDIS will be prepared for voting. 60977 Accelerator performance guideline The comments submitted by national committees were reviewed. Most comments on clause 8.1.6 were deferred until a working party addresses this in 60976. Considerable discussion was held on clause 9.4, on measurements of light field intensity and contrast ratio. Lee Gerig and Geoff Ibbott to review suggestions for periodic measurements, suggest changes (e.g., weekly or daily output checks?) Changes will be reviewed in Delft after which a draft technical report (DTR) will be produced. 60601-2-1, draft 3rd edition. Sethi described the procedure for revising safety standards for compliance with the 3rd edition of the general standard. He and Brian Bibb had prepared a review of the changes needed for the November ’05 meeting. There appear to be two issues: renumbering clauses to align them with the general standard, and adding new provisions required by the 3rd edition. In addition, there are a number of issues that haven’t been addressed by this standard in the past, or at least it is not clear if the existing clauses address the new 3rd edition requirements for “essential performance”, including: SRS/SRT Positioning accuracy Equipment movements that can cause patient injury Accessories Collision avoidance Dosimetry Step-and shoot IMRT Incorrect field shape during delivery Incorrect MU during delivery Record and verify system Neutron doses Whole body dose In-vivo dosimetry Treatment planning system Role in adaptive therapy Electronic imaging devices (e.g., EPIDS) Image orientation Image scale (magnification) and distortion Image detail Field of view and EPID position relative to beam axis Collision avoidance Artifacts Dynamic therapy Position of moving part with respect to MU Collision avoidance Image-guided RT Movement of patient support Movement of MLC Movement of gantry Movement of treatment unit head Dosimetry systems Alternative methods of measuring dose Programmable wedge systems Relationship between position and dose Date of next meeting. The Working group will meet at the time of the next TC meeting October 2-6 in Delft, Holland. Respectfully submitted, Geoffrey S. Ibbott, PhD |
| Issues for AAPM: | AAPM must stay closely involved in the IEC and its working groups. The decisions made by the IEC dictate the design of radiology and radiotherapy equipment, and affect its use in the clinic, as well as issues of radiation protection. These issues will be discussed during a refresher course at the Orlando meeting. |
| Budget Request ($): | Liaison to IEC 2006 Budget Position: Name: Reason(s) for requesting funds: Expenses requested (list and explain): CY 07 CY 07 (AAPM) Travel: 9 members to US TAG meeting @ $500 x 2 mtgs $9,000 3,000 Hotel: 9 members @ $125 x 2 nights x 2 mtgs $4,500 1,500 Per Diem: 9 members @ $40 x 2 days x 2 mtgs $1,440 480 Travel: 1 member @$1,000 x 2 international meetings $2,000 667 Hotel: 1 member @ $125 x 4 nights x 2 mtgs $1,000 333 Per Diem: 1 member @ $40 x 4 days x 2 mtgs $320 107 USNC/ANSI dues: 9 members @ $250 ea. $2250 750 SUBTOTAL $20,510 6837 Each association’s share (1/3 of total) $6,837 6837 Date: July 19, 2006 |

# Text boxes

Liaison to IEC

Dr. Geoffrey S. Ibbott

Funds are requested to support travel to meetings of the US Technical Advisory Group (TAG) to IEC subcommittee 62C, Working Group 1, and to meetings of the Working Group itself.  The AAPM, ASTRO and ACR share these expenses.