| Reported by (Name): | Geoffrey S. Ibbott, Ph.D |
| --- | --- |
| Organization: | International Electrotechnical Commission |
| Position Title: | Chair, Subcommittee 62C. Member, Working Group 1 |
| Activity: | Technical Committee, Subcommittee and Working Group meetings |
| Meeting Dates: | May 7-9, 2007 |
| Meeting Location: | Wiesbaden, Germany |
| Payment $: | Travel reimbursement |
| Reasons for Attending or not Attending | Attended to contribute to the writing of international standards for radiotherapy equipment. |
| Issues from Previous Meetings or Year: | See below. |
| General Description of Activities of the Organization and/or Meeting: | Meeting of IEC Working Group 1 of Subcommittee 62C May 7-9, 2007, Wiesbaden, Germany Technical Committee 62 of the International Electrotechnical Commission deals with electrical equipment used in medicine. It has four subcommittees that deal with (A) general aspects; (B) equipment for imaging; (C) equipment for radiation therapy, nuclear medicine, and dosimetry; and (D) non-radiological equipment. As chair of Subcommittee 62C and a member of its Working Group 1, I am most familiar with the activities of these two groups. This report describes the most recent activities of WG-1. - Working Group 1 of IEC Sub-Committee 62C # Participants The meeting was attended by S. Lillicrap (convenor), H. Sethi and C. Sandin (UK); B. Seidel (secretary), K. Hoeppner, W. Lehmann and W. Goldstein (Germany); I-L. Lamm (Sweden); L. Gerig (Canada); G. Ligthart (Netherlands); A. Marles, S. Johnsen, A. Cohen, S. Coon and G. Ibbott (US). # Agenda The role of the WG is to develop, maintain, and update the standards and technical reports of the IEC that address equipment for radiation therapy. Safety standards that are numbered 60601-2-x deal with medical electrical equipment (defined as equipment that transfers energy to, or receives energy from, or in some way contacts the patient.) These standards modify IEC 60601-1, the so-called General Standard. The 3rd edition of the General Standard was published in December 2005, and incorporated both technical and administrative changes that have required corresponding changes in all of the derivative 60601-2-x standards. In addition, several of the other safety standards and technical reports have come due for revision. Therefore, WG-1 is quite busy. The following standards were reviewed specific details were discussed: - Revision of IEC 60976 This standard establishes performance requirements for linear accelerators used in medicine. It is not a safety standard. The current version has been revised for consistency with the 3rd edition of the general standard, and to accommodate new technologies such as multileaf collimators and on-board imaging. A final draft of this standard is expected to be distributed for voting very soon. - Revision of IEC 60977, 62C/402/CD This is a companion to 60976, and is known as the performance guideline. It contains recommended levels of performance for each of the parameters addressed by 60976. These performance levels are published separately as a technical report to avoid the risk of the recommendations carrying the weight of a standard. A final draft for voting (DTR) was distributed in March for voting by June 1. The US Technical Advisory Group (TAG) has submitted comments and an affirmative vote. - IEC 60601-2-1 Linear accelerator safety standard The accelerator safety standard is being revised substantially for consistency with the 3rd edition of the general standard. At the same time, new clauses are being added to address on-board imaging systems, and the use of MLCs for IMRT, etc. - IEC 60601-2-29, Radiotherapy simulator safety standard Ligtheart had generated a new version of the safety standard to bring it into compliance with the 3rd edition of the general standard. The new version has been distributed to the committee for voting by the end of June. The US TAG has submitted comments and an affirmative vote. - IEC 60601-2-11, Cobalt unit safety standard Goldstein and Sethi had prepared a first draft of a revision of this standard to bring it into compliance with the 3rd edition of the general standard. This revision was complicated by the need to address multi-source cobalt units, such as the Gammaknife. Some consideration was given to the likelihood of development of a combination MR imager/cobalt unit in the future. Goldstein and Sethi will continue work and prepare a committee draft (CD) to be distributed in the near future. - IEC 60601-2-8, X-ray equipment safety standard This standard also will need to be revised to comply with the 3rd edition of the general standard. Gerig and Lamm have agreed to review a first draft prepared by Seidel and provide comments to him. Seidel will distribute a new draft soon afterwards. Ibbott agreed to contact the manufacturers of two electronic brachytherapy systems, Xoft and Zeiss, to determine whether or not this standard applied to their equipment. - IEC 60601-2-17, Remote afterloader safety standard Ibbott had prepared a first draft of a revision for compliance with the 3rd edition, but subsequent changes to the general standard will require further changes to the afterloader standard. Ibbott will make these changes, and also determine if the standard applies to the electronic brachytherapy systems, and distribute a new draft within the next couple of months. - IEC 61217, Coordinates, movements and scales The standard has been amended to address table pitch and roll. A committee draft was distributed for voting by May 18. The US TAG has submitted comments and an affirmative vote. A discussion was held on how to proceed to consider out-of-plane sources, such as are used for x-ray imaging in association with linear accelerators. Hoeppner had prepared a preliminary draft that was discussed in some detail. He will take the suggestions made during the meeting and prepare a new draft. - IEC 62083, Safety of treatment planning systems Ligthart recently distributed a draft revision for compliance with the 3rd edition of the general standard, and an attempt to update the standard to address newer planning systems. Extensive discussion was held on several provisions. The manufacturers will be asked to comment on the latest changes, and Ligthart will prepare a new draft. - Date and venue of next meeting. Marles offered to arrange the next meeting in Greece. The preferred dates are November 12-14. Respectfully submitted, Geoffrey S. Ibbott, PhD |
| Issues for AAPM: | AAPM must stay closely involved in the IEC and its working groups. The decisions made by the IEC dictate the design of radiology and radiotherapy equipment, and affect its use in the clinic, as well as issues of radiation protection. These issues were discussed during a refresher course at the Orlando meeting. |
| Budget Request ($): | Liaison to IEC 2008 Budget Liaison to IEC Position: Dr. Geoffrey S. Ibbott Name: Reason(s) for requesting funds: Funds are requested to support travel to meetings of the US Technical Advisory Group (TAG) to IEC subcommittee 62C, Working Group 1, and to meetings of the Working Group itself. The AAPM, ASTRO and ACR share these expenses. Expenses requested (list and explain): CY 08; CY 08 (AAPM). Travel: 9 members to US TAG meeting @ $500 x 2 mtgs; $9,000; 3,000. Hotel: 9 members @ $125 x 2 nights x 2 mtgs; $4,500; 1,500. Per Diem: 9 members @ $40 x 2 days x 2 mtgs; $1,440; 480. Travel: 1 member @$1,000 x 2 international meetings; $2,000; 667. Hotel: 1 member @ $125 x 4 nights x 2 mtgs; $1,000; 333. Per Diem: 1 member @ $40 x 4 days x 2 mtgs; $320; 107. USNC/ANSI dues: 9 members @ $285 ea.; $2565; 855. SUBTOTAL; $20,825; 6942. Each association’s share (1/3 of total); $6,942; 6942. Date: July 10, 2007 |